MARS 40 Firmware 	Box	50.26.009
		DSP	50.24.009

Stand	65.01.000

Release date: 27 March 2024

Below is a list of things that are new and/or changed in this version compared with 
the previous version.”

Box and stand firmware

1.	The new tribology measuring geometries have been added to the list of compatible accessories.

2.	The new solid clamps for cylindrical samples have been added to the list of compatible accessories.

3.	The new D170/300 pressure cell has been added to the list of compatible accessories.

4.	An initialisation error could occur when a Rheonaut module and a TM-EL-H were connected. This has been solved.

5.	The absolute angle value during oscillation was not correct for values above 2π. This has been solved.

6.	Under certain circumstances the lift motor control did not work correctly when moving the measuring head upward from the measurement position. This has been solved.

7.	When a TM-LI-x module is connected to the device the corresponding valve in the valve block is automatically opened.

8.	Under certain circumstances the instruments display became unreadable after a firmware update. This has been solved.

9.	The Set Element options for Oscillation “Rapid data acquisition Mode” and “Data acquisition during control loop startup phase” now work as intended

-------------------

MARS 40 Firmware 	Box	50.26.008
		DSP	50.24.006

Stand	65.00.041

Release date: 27 November 2023

Below is a list of things that are new and/or changed in this version compared with 
the previous version.”

Box firmware

10.	The rotation CR-Mode control has been improved and now adapts better to the properties of the sample to be measured.

11.	The firmware internal calculation of the rheological properties G’, G”, etc. now checks the plausibility of the measured G’ values, and corrects them if considered useful. This feature is only active when performing measurement at the limits of the possible measuremen range and can be switched on/off by (de)activating the new G’ limiter option in the RheoWin device manager.
Note: For this RheoWin 4.92.0007 or newer is needed.

Stand firmware

12.	Under certain circumstances the lift motor control did not work correctly when moving the measuring head upward from the measurement position. This has been solved.

-------------------

MARS 40 Firmware 	Box	50.25.003
		DSP	50.23.000

Stand	65.00.039

Release date: 01 June 2023

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Note: For the TM-PE-C32 Air and TM-LI-C28 to be recognized by RheoWin, RheoWin 4.91.0011 or newer is needed.

Stand firmware

13.	The firmware is now compatible with the TM-PE-C32 Air temperature module.

14.	The temperature control of the TM-EL-C has been improved.

15.	The temperature control of the TM-PE-P has been improved.

16.	The temperature limits for the TM-LI-C28 have been modified.

-------------------

MARS 40 Firmware 	Box	50.25.003
		DSP	50.23.000

Stand	65.00.017

Release date: 15 september 2022

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.75.0000 or newer. The use of RheoWin 4.86.0002 or newer is recommended.

1.	The firmware is now compatible with the TM-LI-C28 temperature module.

-------------------

MARS 40 Firmware 	Box	50.25.003
		DSP	50.23.000

Stand	65.00.012

Release date: 13 august 2020

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.75.0000 or newer. The use of RheoWin 4.86.0002 or newer is recommended.

1.	The temperature offset tables (in the MTMC driver) can now contain upto 20 values.

2.	The temperature offset tables are now also available for the TM-LI-C, TM-LI-C48, TM-LI-C32 and TM-LI-P modules.

3.	The new TM-PE-P Air (air-cooled Peltier plate module) can now be used with the MARS 40.

4.	The CD-OSC control loop has been further improved.

Please note: For 1. and 2. a new RheoWin MTMC driver (file name MTMC.dll) for RheoWin 4.86.0002 or newer is needed.

Please note: For 3. a hardware update is needed for all MARS 40 build before mid May 2020.

Important: Before updating the firmware save any existing temperature offest table! Updating the firmware will delete any existing temperarature offset table! Either take note of existing offset tables, or export the tables to a *.tot file first. After the firmware update manually enter the table values again, or import the previously exported *.tot file

-------------------

MARS 40 Firmware 	Box	50.25.001
		DSP	50.22.000

Stand	61.09.000

Release date: 09 december 2020

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.75.0000 or newer.

Stand firmware

1) The ejection/release of concentric cylinder geometry rotors has been improved.

-------------------

MARS 40 Firmware 	Box	50.25.001
		DSP	50.22.000

Stand	61.08.000

Release date: 25 february 2019

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.75.0000 or newer.

Stand firmware

1) Under certain circumstances the lift motor control did not work as intended. 
    This has been solved.

-------------------

MARS 40 Firmware 	Box	50.25.001
		DSP	50.22.000

Stand	61.07.000

Release date: 27 july 2018

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.75.0000 or newer.

Stand firmware

1) The maximum lift speed that can be used in measurement elements, mainly in the axial
     ramp element, is now 20 mm/s.

-------------------

MARS 40 Firmware 	Box	50.25.001
		DSP	50.22.000

Stand	61.06.000

Release date: 19 february 2018

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.75.0000 or newer.

Box firmware

1) A bug caused spurious results during the transistion between two CR-ROT elements.
    This has been solved.

-------------------

MARS 40 Firmware 	Box	50.25.000
		DSP	50.22.000

Stand	61.06.000

Release date: 14 november 2017

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.75.0000 or newer.

Stand firmware

1) The intrinsic safety of the lift control has been improved. This does not influence 
    the instrument operation.

-------------------

MARS 40 Firmware 	Box	50.25.000
		DSP	50.22.000

Stand	61.05.000

Release date: 25 september 2017

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.75.0000 or newer.

Box firmware

1) The CD oscillation mode for higher frequencies has been improved.
2) The recognition of a Connect Assist rotor did not work when the RheoWin Monitor 
    (Manual control) window was open when the rotor was mounted. This has been solved.

3) The last set gap value was not always used correctly. This has been solved.

-------------------

MARS 40 Firmware 	Box	50.23.000
		DSP	50.19.000

Stand	61.05.000

Release date: 24 march 2017

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.63.0003 or newer.

Stand firmware

1) The rotor eject function for the situation where the rotor is in measurement position 
    has been improved.
2) The safety mechanism in the lift control for the quick downward movement has been optimised.
3) When the lift ist stopped for safety reasons a message is displayed on the MARS front panel.

-------------------

MARS 40 Firmware 	Box	50.23.000
		DSP	50.19.000

Stand	61.04.000

Release date: 11 january 2017

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.63.0003 or newer.

Box and Stand firmware

1) On some MARS 40 instrument there was a problem with the front display. 
    This has been solved.

-------------------

MARS 40 Firmware 	Box	50.23.000
		DSP	50.19.000

Stand	61.03.000

Release date: 11 november 2016

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.63.0003 or newer.

1) The control of axial ramps has been optimized and is now more accurate.

2) The temperature control of the TM-EL-P has been optimized. In the temperature range of 0 °C to 
    40 ° the set temperature is reached quicker and the stationary temperature value is more stable.

3) The temperature of the Peltier heatsink (of the TM-PE-P and TM-PE-C) is now continuously 
    monitored. When the heatsink temperature reaches an unexpected value (due to insufficient 
    cooling), the Peltier temperature control will be halted and a message will be displayed on the
    instruments control panel.

-------------------

MARS 40 Firmware 	Box	50.22.000
		DSP	50.19.000

Stand	61.02.000

Release date: 10 august 2016

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.63.0000 or newer.

Box firmware

1) Some modifications for production and service personel only have been implemented.

-------------------

MARS 40 Firmware 	Box	50.21.000
		DSP	50.19.000

Stand	61.02.000

Release date: 20 july 2016

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.63.0000 or newer.

Box and Stand firmware

1) The actual lift speed did not correspond to the lift speed values selected in the RheoWin Lift
    element. This has been changed, the actual lift speed now corresponds to the selected value.
    Because the existing selectable lift speed values were relatively low, four new lift speed values
    (5, 10, 20 and 50 mm/s) have been added.

2) The automatic locking of the buttons (apart from the menu button) on the MARS front panel 
    during a Job run, which is useful for preventing accidental lift movement and rotor release when 
    a Job is running, can be now de-activated/activated at any time (i.e. also during a Job run) by 
    using the Menu > Lift (Job) locked / Menu > Lift(Job) free command from the MARS front 
    panel menu. The menu setting is saved when the instrument is switched off and restored when 
    the instrument is switched on.
3) A modification for the torque calibration has been implemented, this is for production and 
    service personel only.

-------------------

MARS 40 Firmware 	Box	50.20.000
		DSP	50.18.000

Stand	61.01.000

Release date: 03 june 2016

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.62.0000 or newer.

Box firmware

1) Improved OSC CD Mode
    The general performance of the deformation (strain) control in oscillation mode (CD-OSC) 
    has been (substantially) improved regarding stability, accuracy, etc.
2) New OSC CD mode with active compliance correction 
    Active compliance correction consists of an active compensation of the compliance of the 
    drive motor + rotor shaft, so that the set deformation value is always achieved. In all 
    previous firmware versions the compliance correction was passive. As long as the 
    deformation is within the linear visco-elastic range, the G’ and G” values resulting from 
    measurements with active and passive compliance correction should be identical.
3) New OSC CS mode with active inertia correction
    Active inertia correction consists of an active compensation of the inertia of the drive motor 
    + rotor shaft, so that the set stress value is always achieved. In all previous firmware 
    versions the inertia correction was passive. As long as the resulting deformation is within 
    the linear visco-elastic range, the G’ and G” values resulting from measurements with active 
    and passive inertia correction should be identical.

-------------------

MARS 40	Firmware 	Box	50.19.000
 		DSP	50.17.000

Stand	61.01.000

Release date: 11 april 2016

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Stand firmware

1) Some modifications were made for compatibility with modified electronics hardware. 
    Firmware version 61.01.000 must be used for all newer MARS 40 (delivered starting from
    2016), using any older firmware in these instrument will make the instrument inoperative.
    Firmware version 61.01.000 is compatible with any MARS 40, that is also with older 
    instruments.

-------------------

MARS 40	Firmware 	Box	50.19.000
 		DSP	50.17.000

Stand	60.04.000

Release date: 08 december 2015

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Box & DSP firmware

1) Under certain circumstances modified Connect Assist adapter rotor properties were not correctly 
    saved into the adapter rotor tag. This has been solved. 
    (This applies to the Connect Assist adapter rotors U1, U2, U3, P1, P2 and ISO only.)

Stand firmware

1) The name of a Connect Assist rotor is now displayed on the MARS control panel with a maximum
    length of 16 (instead of 14) characters, longer names are truncated.

-------------------

MARS 40	Firmware 	Box	50.18.000
 		DSP	50.17.000

Stand	60.03.000

Release date: 28 october 2015

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Stand firmware

1) The normal-force re-zero function has been improved and is more accurate now.

2) The lift-control routines are now compatible with the CTC oven.

(Previous firmware versions were NOT compatible with the CTC oven)

-------------------

MARS 40	Firmware 	Box	50.18.000
 		DSP	50.17.000

Stand	60.02.000

Release date: 22 september 2015

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Box & DSP firmware

1) The buttons on the MARS front panel (apart from the menu button) are now locked when a RheoWin Job is running to prevent accidental lift movement and rotor release. When a locked button is pressed when a RheoWin Job is running, the message "M: Key disabled" will appear on the MARS front panel for a short time.

2) The AutoTension function has been modfied. RheoWin version 4.60.0000 or newer is needed for being able to use the new functionality. See the RheoWin What's new information.

Stand firmware

1) Several modifications for compatibility with certain accessories (RheoScope module etc.) were made.

-------------------

MARS 40	Firmware 	Box	50.17.000
 		DSP	50.17.000

Stand	60.01.000

Release date: 14 july 2015

These are the first release versions for the MARS 60.

----------

Contact Thermo Fisher Scientific

If you have any questions and or suggestions regarding the MARS (firmware)

please send an e-mail to:

support.mc.de@thermofisher.com

----------

This file was last updated on: 27.11.2023 by Jint Nijman.