MARS III Firmware 	Box	50.23.000
 		DSP	50.18.000

Stand	50.25.000

Release date: 09 november 2016

Below is a list of things that are new and/or changed in this version compared with 
the previous version.
Please note: This version is only compatible with RheoWin 4.63.0003 or newer.

1) The actual lift speed did not correspond to the lift speed values selected in the RheoWin Lift
    element. This has been changed, the actual lift speed now corresponds to the selected value.
    Because the existing selectable lift speed values were relatively low, four new lift speed values
    (5, 10, 20 and 50 mm/s) have been added.

2) The control of axial ramps has been optimized and is now more accurate.

3) The temperature of the Peltier heatsink (of the TM-PE-P and TM-PE-C) is now continuously 
    monitored. When the heatsink temperature reaches an unexpected value (due to insufficient 
    cooling), the Peltier temperature control will be halted and a message will be displayed on the
    instruments control panel.

MARS III Firmware 	Box	50.20.000
 		DSP	50.18.000

Stand	50.24.000

Release date: 03 june 2016

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.62.0000 or newer.

Box firmware

1) Improved OSC CD Mode
    The general performance of the deformation (strain) control in oscillation mode (CD-OSC) 
    has been (substantially) improved regarding stability, accuracy, etc.
2) New OSC CD mode with active compliance correction 
    Active compliance correction consists of an active compensation of the compliance of the 
    drive motor + rotor shaft, so that the set deformation value is always achieved. In all 
    previous firmware versions the compliance correction was passive. As long as the 
    deformation is within the linear visco-elastic range, the G’ and G” values resulting from 
    measurements with active and passive compliance correction should be identical.
3) New OSC CS mode with active inertia correction
    Active inertia correction consists of an active compensation of the inertia of the drive motor 
    + rotor shaft, so that the set stress value is always achieved. In all previous firmware 
    versions the inertia correction was passive. As long as the resulting deformation is within 
    the linear visco-elastic range, the G’ and G” values resulting from measurements with active 
    and passive inertia correction should be identical.

MARS III Firmware 	Box	50.17.000
 		DSP	50.17.000

Stand	50.24.000

Release date: 11 april 2016

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Stand firmware

1) Some modifications were made for compatibility with modified electronics hardware. 
    Firmware version 50.24.000 must be used for all newer MARS I (delivered starting from
    2016), using any older firmware in these instruments will make the instrument inoperative.
    Firmware version 50.24.000 is compatible with any MARS I, that is also with older 
    instruments.

MARS III Firmware 	Box	50.17.000
 	DSP	50.17.000

Stand	50.22.000

Release date: 20 november 2015

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Stand firmware

1) The gap zero point is now being searched for until the lowest possible position.

MARS III Firmware 	Box	50.17.000
 	DSP	50.17.000

Stand	50.21.000

Release date: 14 july 2015

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Box & DSP firmware

1) There was a problem with the absolute angle signal during oscillation measurements with f > 1 Hz.
    This has been solved.

2) The quality of the temperature and normal force signals has been improved.

3) The quality of the temperature and gap signals during oscillation measurements has been 
    improved.

4) The Rheonaut Module can now be used together with electrical hood module (TM-EL-H) for 
    improved temperature control at elevated temperatures.

Stand firmware

1) There was a problem with the lift-motor control. This has been solved.
2) The normal force control has been improved.

----------

MARS III Firmware 	RC-Box	50.15.000
 	DSP	50.15.000

M-BoxTC	50.18.000

Release date: 19 feb 2015

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.41.0000 or newer, that

means this firmware version is NOT compatible with older RheoWin versions!

M-BoxTC firmware

1)  The temperature measurement for all temperature modules has been improved.

----------

MARS III Firmware 	RC-Box	50.15.000
 	DSP	50.15.000

M-BoxTC	50.17.000

Release date: 09 july 2014

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.41.0000 or newer, that

means this firmware version is NOT compatible with older RheoWin versions!

M-BoxTC firmware

1)  The temperature control loop for the TM-EL-C has been improved.

----------

MARS III Firmware 	RC-Box	50.15.000
 	DSP	50.15.000

M-BoxTC	50.16.000

Release date: 12 june 2014

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.41.0000 or newer, that

means this firmware version is NOT compatible with older RheoWin versions!

M-BoxTC firmware

1)  There was a bug in the motortemperature measurement. This has been solved.

----------

MARS III Firmware 	RC-Box	50.15.000
 	DSP	50.15.000

M-BoxTC	50.15.000

Release date: 05 august 2013

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.41.0000 or newer, that

means this firmware version is NOT compatible with older RheoWin versions!

RC-Box & DSP firmware

1) The CD-Mode has been further improved.

2) Under certain circumstances in CS-Mode oscillation the actual  number of repetitions 
    was one less as defined in RheoWin. This has been solved.

3) Under certain circumstances the angle reset function (at the start of a measurement) did 
    not work correctly. This has been solved.

4) The firmware now supports upto 60 steps for the RheoWin MultiStep element.

RheoWin 4.41.0000 or newer is needed for this.

M-BoxTC firmware

1)  The firmware now supports three user defined temperature offset tables for special 
     TM-xx-x setups. The three user defined offset tables can be defined and selected in
     the properties of the MTMC in the RheoWin DeviceManager.

RheoWin 4.41.0000 or newer is needed for this.

----------

MARS III Firmware 	RC-Box	50.14.014
 	DSP	50.13.000

M-BoxTC	50.13.000

Release date: 13 december 2012

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Please note: This version is only compatible with RheoWin 4.30.0016 or newer, that

means this firmware version is NOT compatible with RheoWin 4.30.0015 and older versions!

RC-Box & DSP firmware

1) Under very special circumstance the temperature set command did not work properly.
    This has been solved.

----------

MARS III Firmware 	RC-Box	50.14.000
 	DSP	50.13.000

M-BoxTC	50.13.000

Release date: 24 september 2012

Below is a list of things that are new and/or changed in this version compared with 
the previous versions.

This update is recommended to be applied to all MARS III.

Please note: This version is only compatible with RheoWin 4.30.0016 or newer, that

means this firmware version is NOT compatible with RheoWin 4.30.0015 and older versions!

RC-Box & DSP firmware

1) The CD-Mode has been improved in many ways and now gives better, smoother and thus more 
  reproducible results for G’, G”, etc.

2) When the max torque of the drive motor is reached in CD-Mode (during curing for example),
  the firmware will automatically switch to CS-Mode using the highest torque value. 
  The firmware will also switch back to CD-Mode when possible.

----------

MARS III Firmware 	RC-Box	50.13.000
 	DSP	50.12.001

M-BoxTC	50.13.000

Release date: 29 march 2012

Below is a list of things that are new and/or changed in this version compared with 
the previous versions

Please note: This version is only compatible with RheoWin 4.30.0016 or newer, that

means this firmware version is NOT compatible with RheoWin 4.30.0015 and older versions!

RC-Box & DSP firmware

1) A bug which could lead to very rare problems during the instrument initialisation has
    removed.

2) A new torque correction for CS-ROT has been implemented

M-BoxTC firmware

1) The lift control has been modified in such a way that the lift can now always be 
    operated from RheoWin or the instrument front panel by clicking a lift control button or 
    a single button on the instrument front panel. This means that the lift message 
    Manual movement only will not popup anymore when the lift is in a "safety range". 
    This increase in control comfort has been realized by reducing the lift speed under 
    certain special circumstances, for example in the case of a MARS III with a mounted 
    CTC while using a TM-xx-x module.
    This new firmware also solves some other very special lift issues.

----------

MARS III Firmware 	RC-Box	50.11.000
 	DSP	50.11.003

M-BoxTC	50.11.021

Release date: 19 december 2011

Below is a list of things that are new and/or changed in this version compared with 
the previous versions

Please note: This version is only compatible with RheoWin 4.30.0012 or newer, that

means this firmware version is NOT compatible with RheoWin 4.30.0011 and older versions!

M-BoxTC firmware

1) On some (recently build) MARS III units the display did not work properly. This has been solved.
  There is no need to install this firmware version on older MARS III units who do not show this problem.

----------

MARS III Firmware 	RC-Box	50.11.000
 	DSP	50.11.003

M-BoxTC	50.11.003

Release date: 30 november 2011

Below is a list of things that are new and/or changed in this version compared with 
the previous versions

Please note: This version is only compatible with RheoWin 4.30.0012 or newer, that

means this firmware version is NOT compatible with RheoWin 4.30.0011 and older versions!

RC-Box firmware

1) A small problem with the CD-Mode in the previous version was solved. It is strongly advised 
    to update to this version

----------

MARS III Firmware 	RC-Box	50.11.000
 	DSP	50.11.001

M-BoxTC	50.11.003

Release date: 03 november 2011

Below is a list of things that are new and/or changed in this version compared with 
the previous versions

Please note: This version is only compatible with RheoWin 4.30.0012 or newer, that

means this firmware version is NOT compatible with RheoWin 4.30.0011 and older versions!

RC-Box firmware

1) In all previous firmware versions, when using the AutoTension function activated in a
    RheoWin oscillation element, the normal force control was only active in between two
    oscillation periods. This has been changed. 
    The normal force control loop is now working continuously i.e. independ from the oscillation 
    control and measurment.
    This also means that activating the AutoTension function in the Set-Element or in any 
    Osc-Element will result in exactly the same continous normal force control.

2) The optional FastOsc mode (Osc-Time Element, Osc-Temp-Ramp Element) now works
    for any oscillation frequency (previously it worked for f > 1 Hz only).

3) Under certain circumstances the angle signal during CS oscillation was noisy.
    This has been solved.

M-BoxTC firmware
1) Under certain circumstances finding the zero-position did not work correctly due to the

normal force induced by the (higher) lift accelleration. This has been solved.

2) Finding the zero-position with a HS25 measuring geometry did not work.
    This has been solved.

3) The intermediate lift stop position (which can be set using the STOP-Pos. Save 
    command in the menu of the control panel on the front of the instrument) was lost when

the instrument was switched off. This has been solved, the intermediate lift stop position
    is now saved.
4) Under certain circumstances the text display on the control panel was not correct. 
    This has been solved.

----------

MARS III Firmware	RC-Box	50.09.001 
	DSP	50.09.001

M-BoxTC	50.10.000

Release date: 17 july 2011

Below is a list of things that are new and/or changed in this version compared with 
the previous versions

Please note: This version is only compatible with RheoWin 4.30.0001 or newer, that

means this firmware version is NOT compatible with RheoWin 4.30.0000 and older versions!

M-BoxTC firmware

1) On a MARS III with an installed CTC the zero-point could not be reached when a 
    CM-OP-P module (RheoScope) was used together with a TM-EL-H module (electrical
    heated hood). This has been solved.

2) The measuring gap was not set correctly when the options AutoTension and ThermoGap

were both activated in the RheoWin Lift-Element. This has been solved.

----------

MARS III Firmware	RC-Box	50.09.001 
	DSP	50.09.001

M-BoxTC	50.09.000

Release date: 29 april 2011

Below is a list of things that are new and/or changed in this version compared with 
the previous versions

Please note: This version is only compatible with RheoWin 4.30.0001 or newer, that

means this firmware version is NOT compatible with RheoWin 4.30.0000 and older versions!

This update is only needed when using the Rheonaut module with the MARS III.

RC-Box firmware

1) Some issues in the communication between the Rheonaut module controller and 
    the MARS III have been solved.

----------

MARS III Firmware	RC-Box	50.09.000 
	DSP	50.09.001

M-BoxTC	50.09.000

Release date: 14 february 2011

Below is a list of things that are new and/or changed in this version compared with 
the previous versions

Please note: This version is only compatible with RheoWin 4.30.0001 or newer, that

means this firmware version is NOT compatible with RheoWin 4.30.0000 and older versions!

This update is recommended for all MARS III (please check your RheoWin version !!!)

RC-Box and M-BoxTC firmware

1) Lift control has been improved:

- Finding the zero position now takes less time.
  - Finding the zero position after changing from a geometry with a shorter shaft to a geometry 
    with a longer shaft now works automatically.
  - With the new  STOP-Pos. Save command in the menu of the control panel (on the 
    front of the instrument) the current position of the lift is saved as the position at which the 
    lift will be stopped after getting the command to move upwards from RheoWin or from the 
    control panel. By giving the move upwards command again the lift will move to the highest
    possible position.

2) Functionality for the Rheonaut module has been implemented.

3) In FastOsc mode and Rot mode (but not yet in normal Osc mode) the trigger function
    can now provide a block signal or pulsed signal.
    The trigger mode can be set in the MARS III settings in RheoWin, see the RheoWin 
    4.30.0001 readme.doc document.
4) A new Low torque mode which significantly improves the torque resolution in the 
   nanoNewton meter range for oscillation measurements has been implemented.
   For the Low-Torque mode a 2nd MTC table has been implemented to further optimize 
   the performance in the nanoNewton meter range.

The Low-Torque mode can be activated in certain RheoWin oscillations measuring 
   elements, see the RheoWin 4.30.000x readme.doc document.
   Please note: The full effect of the Low torque improvements is only achieved with 
   instruments which were delivered with this firmware version (or a newer version) from 
   the factory.

5) A problem with the initalisation of the AC-amplifier (in a very limited number of 
   MARS III instruments) has been fixed.

----------

MARS III Firmware	RC-Box	50.06.000 
	DSP	50.07.001

Release date: 22 october 2010

Below is a list of things that are new and/or changed in this version compared with 
version DSP 50.06.000

Please note: This version is only compatible with RheoWin 4.20.0000 or newer, that

means this firmware version is NOT compatible with RheoWin 4.10.00xx and older versions!

RC-Box DSP firmware

1) Under certain circumstances there were problems when running long duration oscillation 
    measurements. This has been solved.
    This update is recommended for all MARS III (please check your RheoWin version !!!)

----------

MARS III Firmware	RC-Box	50.06.000 
	DSP	50.06.000

M-BoxTC	50.06.000

Release date: 03 july 2010

Below is a list of things that are new and/or changed in this version compared with 
version RC-Box 50.04.000, DSP 50.02.000, M-BoxTC 50.05.005.

Please note: This version is only compatible with RheoWin 4.20.0000 or newer, that

means this firmware version is NOT compatible with RheoWin 4.10.00xx and older versions!

RC-Box and M-BoxTC firmware

1) The new TM-PE-C Peltier cylinder temperature module is now supported by the firmware.
    When the TM-PE-C is connected to the MARS III the display will show the text PE-C.

2) The display on the MARS III now displays warning messages when a temperature module
    is not connected correctly.

3) The temperature control loop of the TM-EL-P module has been improved.

4) There was a problem with the temperature offset tables, this has been solved.

5) The new FastOsc method has been implemented in the firmware. With this new method 
    a data acquisition rate of one rheological data point per 2 ms in oscillation can be 
    achieved. This is usefull for fast UV-curing measurements for example.

6) The CD-Mode control loop for the solid clamps has been optimized.

7) Lift control has been optimized. The lift will now always stop when the stop button on
    the display is pressed.

8) The determination of the normal force zero point has been improved.
9) The TCP/IP communication between the MARS III and RheoWin has been optimized.

10) The new UTMC control box (which is needed for the MARS II and RSxxx when using 
    the TM-XX-Y temperature modules) is now supported by the MARS III. 
    Please note that the UTMC control box is (normally) NOT needed for a MARS III !!

----------

MARS III Firmware	RC-Box	50.04.000
	DSP	50.02.000

Release date: 19 january 2010

Below is a list of things that are new and/or changed in this version compared with 
version RC-Box 50.03.000, DSP 50.01.000.

Please note: This version is only compatible with RheoWin 4.10.0002 or newer, that

means this firmware version is NOT compatible with RheoWin 4.10.0001 and older versions!

Very important: When the current firmware version in your MARS III is version 50.03.000 (and only in that case) you MUST switch the MARS III off and on again after putting it into update mode, the rest of the firmware update is the same as before. This is because of a bug in firmware version 50.03.000. If you don’t switch the MARS off and on again after putting it in update mode the firmware updater will not be able to detect the MARS III and the new firmware can not be uploaded. We apologize for the inconvenience.

RC-Box Firmware

1)  With the RC-Box firmware version 50.03.000 a firmware update is only possible when 
     the instrument is switched off and on again after putting it into update mode.
     This has been solved.

2)  Under certain circumstances the instrument would never finish the initialisation.
     This has been solved.

----------

MARS III Firmware	RC-Box	50.03.000
	DSP	50.01.002

M-BoxTC	50.05.000

Release date: 9 december 2009

Below is a list of things that are new and/or changed in this version compared with 
version RC-Box 50.02.000, DSP 50.01.002, M-BoxTC 50.04.000.

Please note: This version is only compatible with RheoWin 4.10.0002 or newer, that

means this firmware version is NOT compatible with RheoWin 4.10.0001 and older versions!

RC-Box and M-BoxTC firmware

1) The trigger relais was not properly controlled. 
    This has been solved.

2) The measurement data of an external signal showed intermittant missing values. 
    This has been solved

3) The optimize function for the number of repitions for oscillatory measurements

produced wrong values for the phase angle.

This has been solved.

4) The instrument initialisation procedure itself and the display of information during the 
    initialisation has been optimized.

----------

MARS III Firmware	RC-Box	50.02.000 
	DSP	50.01.000

M-BoxTC	50.04.000

Release date: 16 october 2009

Below is a list of things that are new and/or changed in this version compared with 
version RC-Box 50.01.000, DSP 50.01.000, M-BoxTC 50.01.000.

Please note: This version is only compatible with RheoWin 4.10.0002 or newer, that

means this firmware version is NOT compatible with RheoWin 4.10.0001 and older versions!

RC-Box Firmware

1)   New functionality for the fluid valve control display in the RheoWin MTMC 
      diagnosis dialog has been implemented.

2)   The minimal and maximal allowed values for the normal force control were not correct,  this has been corrected.

M-BoxTC Firmware

1)   The safety functions for the temperature sensor have been optimized

2)	The TM-xx-C module recognition has been modified.

3)   The temparature control for the TM-xx-C modules has been optimized.

4)   Control functionality for the external power supplies for the TM-xx-x modules has been implemented.

5)   New functionality for the fluid valve control display in the RheoWin MTMC 
      diagnosis dialog has been implemented.

6)   New display and lift functionality for the CTC oven has been implemented.

7)   The display functionality for finding the zero point and setting the gap has been improved. 
      Directly after selecting the command from the display menu, the display will now switch 
      back to the default menu so that the gap change can be observed.

----------

Contact Thermo Fisher Scientific

If you have any questions and or suggestions regarding the MARS (firmware)

please send an e-mail to:

support.mc.de@thermofisher.com

----------

This file was last updated on: 28.06.2016 by Jint Nijman.