MARS iQ	Firmware 	V1 	01.18.042
 			V2		01.18.022

V3 	01.18.012

HMI 	01.18.000
Release date: 27 February 2024

Below is a list of things that are new and/or changed in this version compared with the previous version.

Firmware V1, V2, V3 and HMI

1.	The γ (deformation) value measured and displayed (on the touchscreen and in RheoWin) during oscillation measurements was not correct in the previous firmware version only. This has been solved.
Note: The value of all other oscillation data (G’ G” etc.) was not affected by this!.

2.	The data acquisition during a ROT ramp in CS-Mode was improved.

3.	The control of the measurement gap (during a measurement) has been improved.

4.	Under certain circumstances the “maximum normal force error message” was not cleared correctly and was displayed at a jo start. This has been solved.

5.	Under certain circumstances reading temperature offset values from the TM-CR-O control box did not work correctly. This has been solved.

6.	The handling of the TM-CR-O “oven not closed” message was improved.

7.	The max. allowed motor temperature was increased for use at ambient temperatures up to 35 °C.

8.	The max. cooling power of the TM-PE-P Air for T > 180 °C was slightly reduced due to safety issues.

9.	At ambient temperatures above 26 °C (which are not recommended) the max. heating power of the TM-EL-P + TM-EL-H combination was slightly reduced.

----------

MARS iQ	Firmware 	V1 	01.18.008
 			V2		01.18.010

V3 	01.18.001

HMI 	01.18.000
Release date: 10 January 2024

Below is a list of things that are new and/or changed in this version compared with the previous version.

Firmware V1, V2, V3 and HMI

1.	Under certain circumstances a time-out in the communication (for example due to certain windows power saving options) between the instrument and RheoWin would lock up the instrument. This has been solved.

2.	The CD-OSC control loop can now (optionally) switch to CS mode when the torque is below a certain value. RheoWin 4.92.0000 or higher is needed for using this option.

3.	The MARS iQ is now compatible with the new TM-CR-0450 oven.

4.	When the air pressure for the airbearing is above 3.0 bar the instrument will now display an error message.

5.	The MARS iQ is now compatible with the new Tribology measuring geometries. 
Note: For tribology measurements the air pressure must be set to 2.5 bar. The instrument will display an error message when this is not the case.

6.	Under certain circumstances not all data points, to be acquired as part an oscillation measurement, were actually transfered to RheoWin and saved. This has been solved.

7.	The instruments touchscreen now always displays the (uncorrected) motor torque and motor angle values.

8.	The temperature control of the TM-PE-x Air modules (around ambient temperature) has been improved.

9.	 The functionality of the LOCK button has been enhanced. When touching the LOCK button for more then 3 seconds a menu will be displayed which allows the operator to select the current angular position as the lock position.

10.	Some display text in Italian (and a few other) language(s) were not displayed correctly. This has been solved.

11.	The autotension control loop has been improved.

12.	The CR ROT control loop has been improved.

13.	The MARS iQ can now control an external gas valve for the air-bearing air/gas supply. 
  This is meant for use in a glove box when using inert gas for the air-bearing.

----------

MARS iQ 	Firmware 	V1 	01.17.000
 			V2		01.15.000

V3 	01.17.000

HMI 	01.17.00
Release date: 16 June 2023

Below is a list of things that are new and/or changed in this version compared with the previous version.

Firmware V1, V3 and HMI

1.	The new tribology rotors are now recognized by the instrument.

2.	The AutoTension control loop has been improved.

3.	The error handling for the normal force sensor has been improved.

4.	The error handling for the temperature control has been improved.

----------

MARS iQ	Firmware 	V1 	01.16.000
 			V2		01.15.000

V3 	01.15.000

HMI 	01.04.00
Release date: 26 April 2023

Below is a list of things that are new and/or changed in this version compared with the previous version.

Firmware V1, V2, V3 and HMI

5.	The upper temperature alarm limits for the individual TM-xx-x have been modified and are now set to 2 °C above the maximum specified temperatures.

6.	The new TM-PE-C32 Air is now recognized and controlled by the instrument.

7.	The control of the TM-PE-x and TM-PE-x Air modules has been improved.

8.	A possible overheating of the TM-PE-x modules due to a lack of counter-cooling is now recognized and stopped earlier and more effective.

9.	The MARS iQ firmware now includes two operation-hour counters. One for counting the power-on time and one for counting the actual measurement time. The counter values are displayed in the Configuration > Device info menu on the instruments touch screen. Starting from RheoWin 4.91.0010 these values are stored in every RheoWin *.rwd data file.
NOTE: Of course both counters start with zero hours from the moment this new firmware is installed.

10.	After using the Lock button on the Lift page of the touchscreen user interface the geometry recognition did not work correctly anymore. This has been solved.

11.	There was a problem with the lift zero-point for certain geometries that are based on a ConnectAssist adapter. This has been solved.

12.	The lift speed during a Job is not limited anymore when the measuring geometry (rotor) was not correctly recognized.

13.	The CR-Mode control loop has been improved.

----------

MARS iQ	Firmware 	V1 	01.14.000
 			V2		01.10.024

V3		01.12.000

HMI 	01.03.09
Release date: 30 January 2023

Below is a list of things that are new and/or changed in this version compared with the previous version.

Firmware V1

14.	Under certain circumstances updating the firmware (with version 01.12.000) did not work correctly. This has been solved for this version.

15.	Under certain circumstances a erroneous problem with the Peltier heat exchanger temperature was reported. This has been solved.

16.	An unnecessary pause in the execution of an axial ramp has been removed.

17.	The bearing friction measurement which is part of the instruments initalisation process has been improved.

----------

MARS iQ	Firmware 	V1		01.12.000
 			V2		01.10.024

V3		01.12.000

HMI 	01.03.09
Release date: 30 August 2022

Below is a list of things that are new and/or changed in this version compared with the previous version.

Firmware V1, V2, V3 and HMI

1.	The acoustic signal which alerts the operator to remove any rotor from the drive motor shaft (in order to be able to continue the initialisation of the instrument) has been modified. It is now active until the rotor is removed.

2.	The initialisation of the valves in the valve block has been optimized.

3.	Under certain (very rare) circumstances a problem in the temperature control electronics could lead to an unexpected heating-up of a TM-xx-x. This has been solved.

4.	Under certain circumstances the touchscreen would not work as intended after waking the instrument up from the standby mode. This has been solved.

5.	Setting Temperature Offset values from RheoWin did not work correctly. This has been solved.

6.	The new TM-LI-C28 is now recognized by the instrument.

----------

MARS iQ	Firmware 	V1		01.10.021
 			V2		01.10.024

V3		01.10.004

HMI	01.03.09
Release date: 15 February 2022

Below is a list of things that are new and/or changed in this version compared with the previous version.

Firmware V1, V2, V3 and HMI

1.	Under certain circumstances starting a Job from the MARS iQ touchscreen did not work correctly.
This has been solved.
Note: Further changes regarding this functionality have been made in RheoWin 4.88.0000 and newer.

2.	A screensaver for the touchscreen has been implement. The screen saver becomes active after not using the touchscreen for 30 minutes. To stop the screensaver just tap the touchscreen.

3.	Under certain circumstances the OSC Temperature Stepped Ramp element did not work as intended. This has been solved.

4.	When the max. torque or the max. speed is reached during a ROT measurement, the active element in RheoWin will be stopped and the Job will continue with the next element.

5.	When the max. torque is reached during an OSC frequency or amplitude sweep, the active element in RheoWin will be stopped and the Job will continue with the next element.

6.	Under certain circumstances the CS OSC Time curve element did not work as intended. This has been solved.

7.	The sample volume needed for the attached rotor (measuring system) is now displayed on the touchscreen (in the geometry status information menu).

8.	Other minor bugs have been solved.

----------

MARS iQ	Firmware 	V1	 01.09.007
 			V2	 01.09.005

V3	 01.09.000
Release date: 05 August 2021

Below is a list of things that are new and/or changed in this version compared with the previous version.

Firmware V1

1.	The tolerance for the friction determination which is part of the initialization routine was modified.

----------

MARS iQ	Firmware 	V1	 01.09.000
 			V2	 01.09.005

V3	 01.09.000
Release date: 20 April 2021

Below is a list of things that are new and/or changed in this version compared with the previous version.

Firmware V1, V2, V3

1.	Initialization issues are now displayed on the instrument’s touchscreen in a more understandable way.

2.	Under certain circumstance the automatic geometry recognition did not work correctly. This has been solved.

3.	Under certain circumstances the torque and angular velocity values measured during a CR ramp were suboptimal. This has been improved.

4.	Network IP address configuration:
- Changing the static IP address now takes immediate effect (that is, the instrument does not need to be switched off/on for the new IP address to be applied).
- The instrument will now automatically be assigned an IP address by a network DHCP server when the instrument is switched on.
- The instrument will now automatically be assigned an IP address by a network DHCP server when the instrument is connected to a network when the instrument is already up and running.

5.	When a TM-LI-X temperature control module is connected to instrument, the corresponding (lower) valve in the valve block is automatically opened completely. 
Note: A TM-LI-X module must only be connected to the valve block when the temperature of the circulating fluid is in the range of -25 °C < T < 60 °C. See the MARS TM-XX-X instruction manual.
Note: When disconnecting a TM-LI-X from the rheometer first unplug the electrical connection and give the instruments a few seconds time to close the valve automatically.

6.	The display of information on the instrument’s touchscreen when a RheoWin Job is running has been improved for several Job elements.
Note: RheoWin 4.87.0010 must be used for this to take full effect.

7.	In the configuration menu on instrument’s touchscreen there is a new menu item Corrections. Tapping the menu item will open the Corrections dialog. From that dialog the instruments bearing friction calibration routine and the instrument torque mapping calibration routine can be started.

8.	The temperature set dialog can not be accessed anymore when no TM-XX-X or a TM-LI-X is connected the instrument.

----------

MARS iQ	Firmware 	V1	 01.07.000
 			V2	 01.06.002

V3	 01.05.001

Release date: 25 November 2020

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Firmware V1

1.	The lift control for the combination of a TM-xx-C (C=Cylinder) temperature module plus the TMP adapter and a plate or cone rotor (upper part of the measuring geometry) has been improved.

----------

MARS iQ	Firmware 	V1	 01.06.003
 			V2	 01.06.002

V3	 01.05.001

Release date: 13 August 2020

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Firmware V1, V2, V3

1.	The temperature offset tables (in the MTMC driver) can now contain upto 20 values.

2.	The temperature offset tables are now also available for the TM-LI-C, TM-LI-C48, TM-LI-C32 and TM-LI-P modules.

3.	Lift control for the following measuring geometries was improved: Interfacial Ring rotor, interfacial double bi-cone rotor, ISO (Brookfield) spindles, the universal adapters U1, U2, and U3, the UV upper plate and the pressure cell magnetic coupler.

4.	The TM-EL-C module is now always recognized correctly.

Please note: For 1. and 2. a new RheoWin MTMC driver (file name MARSIQT.dll) is needed.

Important: Before updating the firmware save any existing temperature offest table! Updating the firmware will delete any existing temperarature offset table! Either take note of existing offset tables, or export the tables to a *.tot file first. After the firmware update manually enter the table values again, or import the previously exported *.tot file

----------

MARS iQ	Firmware 	V1	 01.05.001
 			V2	 01.06.000

V3	 01.05.000

Release date: 14 May 2020

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Firmware V1, V2, V3

1.	The OSC frequency sweep and Osc amplitude sweep now work with tabular set values for the frequency resp. amplitude. (With RheoWin version 4.86.0017 or higher).

2.	The ROT Multistep element has been implemented.

3.	The CD-OSC control loop has been further improved.

4.	Normal force control has been improved.

5.	The raw phase angle was missing in the OSC data. This has been solved.

6.	The accuracy of the lift positioning for intermediate positions has been improved.

7.	The ConnectAssist geometry recognition functionality has been optimized.

8.	Under certain circumstances the Lift control did not work properly. That has been solved.

9.	The MARS iQ (Air) is now compatible with the new TM-P-PE Air (air cooled Peltier plate module).

10.	The MARS iQ (Air) is now compatible with the new Protect Assist sample area shield door.

Touchscreen user interface (HMI)

1.	A splash screen is displayed for a few seconds after switching on.

2.	The initialisation screen has been enhanced and now contains a progress bar.

3.	The geometry and temperature control unit information screens can now be accessed directly from the home screen by tapping on the corresponding area of the status bar.

4.	The temperature can now be controlled from a popup dialog which is opened by tapping on the temperature area in the statusbar. A temperature value can be entered, temperature control can be started and stopped.

5.	The home screen now has an android style Configuration menu with the following menu items: Temperature control, Network, Start menu, Language, Lift.

6.	On the Temperature control men an initial start temperature value be set.

7.	On the Network menu the instruments IP adress can be set.

8.	On the Startmenu menu either the Home screen or the Lift control screen or the Job launch screen can be selected as the default start menu.

9.	On the Language menu the language for the touchscreen user interface can be selected. Currently 18 languages are available: English, German, Dutch, French, Spanish, Italian, Portugese, Polish, Czech, Slovakian, Finnish, Russian, Turkish, Chinese, Japanese, Korean, Thai, Indonesian.

10.	On the Configuration Lift menu the lift speed for the slow up/down lift button can be set.

11.	The current lift position can be set as the zero position by tapping the lift zero button for more than 3 seconds.

12.	A lower stop position for the lift can be set (in a popup dialog) by tapping the fast down lift button for more than 3 seconds.

13.	The Lift screen now displays a clear message when the lift is moving slower as expected.

14.	The Job Launch list is now numbered.

15.	The temperature control information screen shows the status of the upper and/or lower valve in the valve block or the status of the TM-PE-P Air fan if applicable.

16.	The range bars for torque, angular velocity and angle have been improved.

----------

MARS iQ	Firmware 	V1	 01.03.000
 			V2	 01.02.000

V3	 01.03.000

Release date: 18 december 2019

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Firmware V1, V2, V3

1) The logaritmic data point distribution of an OSC Amplitude Sweep element was modified.
    It is now identical to that of a MARS 40/60.

2) The ROT Multistep element has been implemented.

3) The new TM-PE-P Air is now recognized.

4) The motor/bearing friction correction has been slightly modified.

5) The number of repetitions and number of waiting periods for all OSC elements are now
     user definable (in RheoWin).

6) Certain internal lift functionality has been optimized.

7) The ConnectAssist geometry recognition functionality has been optimized.

8) The instrument type indentification was not completely correct.

This has been solved.

9) Touching an "empty" space on the touchscreen lead to an unintended reaction. 
    This has been solved.

----------

MARS iQ	Firmware 	V1	 01.01.000
 			V2	 01.01.001

V3	 01.01.002

Release date: 30 july 2019

These are the first release versions of the MARS iQ firmware.

----------

Contact Thermo Fisher Scientific

If you have any questions and or suggestions regarding the MARS (firmware)

please send an e-mail to:

support.mc.de@thermofisher.com

----------

This file was last updated on: 27.02.2024 by Jint Nijman.