VT iQ	Firmware 	V1	01.10.004
 		V2	01.02.002
		HMI	01.01.03

Release date: 16 october 2023

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Firmware V1

- The firmware now monitors the counter cooling of the Peltier element in the TM-PE-x modules and
  displays a warning message when the cooling is not sufficient (most likely because of problems 
  with the hoses between the TM-PE-x module and the HX iQ).

- There was a problem with the maximum number of Job elements (10 elements for an intermal 
  job only). This has been solved.

- The display of the range bars during a measurement has been improved.

------------

VT iQ	Firmware 	V1	01.09.000
 		V2	01.02.002
		HMI	01.01.03

Release date: 26 june 2023

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

Firmware V1 and V2

- Sending temperature offset tables from RheoWin to VT iQ (Air) did not work correctly. 
  This has been solved.
- The external Pt100 sensor was not correctly recognized. This has been solved.

- The PC170 (pressure cell) was not correctly recognized. This has been solved.

- The acquisition of OSC data has been improved. The relation between the number of waiting 
  periods and number of measurement periods was changed, the sum of these two numbers 
  stays the same.

------------

VT iQ	Firmware 	V1	01.08.000
 		V2	01.01.002
		HMI	01.01.03

Release date: 14 april 2023

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- The VT iQ firmware now includes two operation-hour counters. One for counting the power-on time and one for counting the actual measurement time. The counter values are displayed in the Configuration > Device info menu on the instruments touch screen. Starting from RheoWin 4.91.0010 these values are stored in every RheoWin *.rwd data file.
NOTE: Of course both counters start with zero hours from the moment this new firmware is installed.

------------

VT iQ	Firmware 	V1	01.06.000
 		V2	01.01.002
		HMI	01.01.03

Release date: 06 april 2023

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- When the gap setting for a parallel plate geometry was different from the default value the M-factor and thus the shear-rate was not calculated correctly. This only applies to measurement started from the instruments touch screen, that is, NOT to measurements started from RheoWin. This has been solved.

------------

VT iQ	Firmware 	V1	01.04.000
 		V2	01.01.002
		HMI	01.01.03

Release date: 30 january 2023

Below is a list of things that are new and/or changed in this version compared with 
the previous version.

- The temperature offsets for the TM-PE-C in combination with a cylinder measuring geometry
   were not working correctly. This has been solved.

- The (optional) notification beep at the end of a Job was not always terminated correctly. 
   This has been solved.
- The numerical results of the evaluation elements are now displayed in floating point format 
    (that is not in scientific format anymore).
- During a firmware update a message is now displayed on the instruments touch screen.

------------

VT iQ	Firmware 	V1	01.02.000
		V2	01.01.002

HMI	01.01.03

Release date: 21 september 2022

These are the first release versions for the (redesigned) VT iQ build after June 2022.

These instruments have 3 USB ports on the left hand side of the instrument head and one round connector for connecting a TM-xx-x temperature module plus a separate D-Sub connector for the HX iQ heat exchanger

----------

Contact Thermo Fisher Scientific

If you have any questions and or suggestions regarding the MARS (firmware)

please send an e-mail to:

support.mc.de@thermofisher.com

----------

This file was last updated on: 16.10.2023 by Jint Nijman.