What's new in RheoWin version 4.50.0017

(Released 19.02.2015)

Below is a list of things that are new and/or changed in version 4.50.0017 compared with 
version 4.50.0013 of RheoWin.

- DeviceManager

- Viscotester iQ 
    - The service/calibration date is now only checked for firmware versions 01.04.020 and higher

because it caused problems with earlier firmware versions.

- The temperature value displayed in the Device toolbar was not correct during oscillation
      measurements. This has been solved.

- Image acquisition

- Under certain circumstances there were problems with the image acquisition routines when 
      testing the camera before setting the camera properties in the DeviceManager. 
      This has been solved.

- JobManager

- Snapshot

- There was a problem when running a Snapshot test. This has been solved.

- Show data and wait Element
    - In the Show data and wait element the communication with a temperature control unit is
      now deactivated also when the communication with the device is deactivated.

- All OSC Elements

- Under certain circumstances the editors of these elements could crash. 
      This has been solved.

- Curvefit Element

- Under certain circumstances the parameter values displayed in the information window were not

correct when the physical unit of the parameters were displayed together with the parameters.
      This has been solved.

- In the Numerical group box on the Layout page of the element editor there are three new 
      entries Y value at X = with which the Y value of the curve-fit model can be calculated for 
      three X values. The result is displayed in the information window.
  - Crossover Element
    - The time t, the segment time t_seg and the temperature T are now available as QC criteria.

- DataManager

- Saving images

- Camera images can now additionally be saved as a video file in the WMV format by right
      clicking on the image area and selecting the Save as video command.

Contact Thermo Fisher Scientific

If you have any questions and or suggestion regarding RheoWin

please send an e-mail to:

support.mc.de@thermofisher.com

----------

This file was last updated on: 19.02.2015 by Jint Nijman.